## Cover Page
GENERAL g ELECTRIC
Doc ID: | DOC2070884
Design Verification Test Procedure
Project Name: Venue
Module Name: Venue R2 M3.1 PS-SW Monitoring and Error Handling VPR
Revision History:
Rev # | Rev. Date | Name | Changes Description
Rev 1.0 | 2018-01-10 | Alesia Azarov | Initial release
Rev 2.0 | 2018-01-29 | Alesia Azarov | Test ID 6251 has been removed Adding tests 6123, 6124 and 8000
Rev 3.0 | 2018-06-20 | Kostia Ivanzov | Test ID 6264 was updated according to requirement.
Rev 4.0 | 2019-07-14 | Nika Goferman | Following test IDs added: 1618, 1692, 1741, 1743, 1793, 1798, 1761, 2432, 11731, 1739, 11735. Following test IDs removed: 7880, 8874, 6263. Modified Tests IDs: 6255, 6252, 6311, 6343, 5742, 5747, 6283, 6562
Rev 5.0 | 2019-07-30 | Ami Peled | Test IDs 6253, 6246, 6277 updated
| Name |  | Function | Signature | Date
Author | Ami Peled |  | Verification Engineer | See My WorkShop
Reviewer | Yulia Sne |  | V&V Manager
Approver | Sofi Novofastovski |  | LSD
* This doc extracted from HP-ALM

## Intro
1 |  | Introduction
| 1.1 | Scope
This document's ( DOC2070884) purpose is to test the Venue R2 M3.1 PS-SW Monitoring and Error Handling VPR of the Venue system, according to the system's Product Specification (DOC2058182).
| 1.2 | Test Sets
This procedure divided to the following test sets:
|  | Test Set 1 | Battery
|  | Test Set 2 | Burning
|  | Test Set 3 | Display
|  | Test Set 4 | Labeling
|  | Test Set 5 | Preventive Actions
|  | Test Set 6 | Probes failure handling
|  | Test Set 7 | SW Warnings & Notifications
|  | Test Set 8 | Thermal Control
2 |  | Equipment Setup
The following equipment is required. The calibration schedule must be checked prior to testing for each simulator, function generator, oscilloscope, and multimeter. Attach an additional equipment list to the report, if required.
3 |  | Pass Criteria:
Pass criteria will be defined in the test verification summary report

## Summary
# | Test Name | Tested Subject
1 | 02 - Battery life threshold | 1. Verifyting that when the capacity of a full-charged battery is below 50% of a new battery capacity, then the Battery Health indication appears. 2. Verifying that when the capacity of a full-charged battery is above 70% of a new battery capacity, then the Battery Health indication does not appear.
2 | 07 - MPB - System shut down due to overheat - Upper Limit(For Trident) | Setting one of the sensors associated with the MPB fan, above "Upper Limit" for more than 6 minutes.
3 | 04 - BE - System shutdown due to overheat - Upper Limit(For Trident) | 1. Setting one of the sensors associated with the BEP fan, above"Upper Limit" for more than 6 minutes.
4 | 01 - Battery capacity display | 1.Verifying that the system has a battery indicator 2.Verifying that the user is notified when the remaining system battery capacity is less than 10 min.
5 | 05 - Error indicator of disabled battery charge | Verifying that a special error indication appears when AC plug is connected to power but charging of battery is disabled
6 | 02 - Low battery pop up message | 1. Verifying that a pop up messsage about low battery capacity apperas when a remaining time is 10 min. 2. Verifying that a pop up messsage about low battery capacity re-apperas every 2 minutes after tapping ok on the message. 3. Verifying that the system automaticaly trigers "End Exam" and turns Off the system when remaining time is 2 min.
7 | 06 - Special indication when battery scanning time is less than 10 min | Verifying that battery indicator (auxiliary display) has a special indication when remaining scan time is less than 10 min.
8 | CFE GEN2 FW burning via gficonsole tool | The ability to update FE Monitoring FW version via gficonsole tool
9 | FW reading from EEPROM | Verifying each board details are read and presented correctly under Config
10 | 01 - Auto adjustment of display brightness | Auto adjustment of the monitor's brightness .
11 | 04 - MI limitation for Opthalmic preset | 1. Verifying that the Mechanical Index (MI) is less than 0.23 for optalmic preset in 12L-RS probe 2. Verifying that the Mechanical Index (MI) is less than 0.23 for optalmic preset in 12Ln-RS probe 3. Verifying that the Mechanical Index (MI) is less than 0.23 for optalmic preset in 3Sc-RS probe 4. Verifying that the Mechanical Index (MI) is less than 0.23 for optalmic preset in 9L-RS probe
12 | 07 - The factory name of preset Ophthalmic | 1. The name of the Ophthalmic factory preset on L12n-RS probe 2. The name of the Ophthalmic factory preset on 12L-RS probe 3. The name of the Ophthalmic factory preset on 3Sc-RS probe
13 | 09- Labeling - Service Manuals | The Service Manual provides information for service personnel regarding: 1. FRU replacements 2. Installations 3. Maintenance 4. Quality checks
14 | 01 - Preventation of scanning in case of over voltage | 1. Verifying that the a probe is scanning with in limit transmit voltages 2. Verifying that the probe stops scanning in case of transmission is over voltage
15 | 01 - Scanning relation to FE clock frequency | Verifying that the Image degradates due to HW error
16 | 01 - Periodic system shutdown mechanism | 1. Verifying that user is able to postpone periodic shutdown 2. Verifying that periodic shutdown mechanism is done upon user aproval
17 | 01 - Probe ID comparison in the system | Verifying that supported probe is not supported if Probe ID in system file was changed
18 | 02 - Probe not supported if probe ID is duplicated | Verifying that supported probe is not supported ifprobe ID is duplicated in the system
19 | 01 - Presence of Thermal and Mechanical indexes | Verifying that Thermal and Mechanical indexes are presented
20 | 02 - Correct presentation of geometry data | Verifying that geometry data represented correctly by using tissue phantom
21 | 03 - Correct representation of colour coded data | Verifying that colour coded data represented correctly by using flow phantom
22 | 06 - Confirmation of patient data assignment | Verifying that the user must confirm intended assignment of unassigned patient
23 | 07 - Includement specific risks per tool in user manual | Verifying that the user manual includs specific risks for the following automatic M&A tools: 1. Auto-VTI 2. Auto-IVC 3. B-Lines
24 | 11 - M&A graphics position | Verifying that the M&A is performed, remove the measurements from the image upon changing the layout of a combined image
25 | 12 - Discontinuities in time motion display | Verifying that moving TM-cursor results in recordings with discontinuites are indicated on the display
26 | 13 - Colour flow presentation | Verifying that colour flow processing on flow phanton represented correctly
27 | 14 - M&A identifiers | Verifying that M&A identifires appear aside each measurement in resault screen
28 | 16 - Confirmation of data deletion | 1. Tapping on "No" in deletion confirmation window 2. Tapping on "Yes" in deletion confirmation window
29 | 28 - Real-time Clock | 1. The real-time clock indicates the current time in the Live scanning modes 2. The real-time clock indicates the current time in the Freeze state
30 | 17 - Patient data replacement message | Verifying that a warning massage that a patient data will be replaced with the backup copy appears
31 | 27 - Keep or replace message when storing to media | Verifying that the message asking to keep or replace already existing data on media while storing to media appears
32 | 18 - No space on media report | Verifying that copying images to a full media not performed and failure reported
33 | 19 - Network monitoring and patient demographic backup | 1. Verifying that network is monitored 2. Verifying that patient demographic data is backed up on the system 3. Verifying that patient data is backed up on the Local archive
34 | 20 - SW version and HW configuration on a configuration info screen | 1. Verifying that SW version appears on configuration information screen 2. Verifying that HW configuration appears on configuration information screen
35 | 21 - Rejection of image in spooler | Verifying that storage job spooler notifies about images that were rejected
36 | 22 - Probe visualisation by the system | a) Verifying that the system provides probe menu during selection b) Verifying that the system provides probe display during operation
37 | 23 - Preventation of scanning if HW not valid | Verifying that scanning is prevented if HW cofiguration validation failed
38 | 24 - Control for Image Quality optimisation | Verifying that the system provides a control to activate test-pattern for Image Quality optimisation
39 | 02 - message about light level | message appears when light level exceeds the recommended level
40 | 02 - Temperature and Fan speed Periodic Monitoring | 1. Verifying that the log contains periodic readings of temperatures sensors from the following components: a) FE b) BEP c) Batteries 2. Verifying that the log contains periodic readings of the systems fans.
41 | 01 - Fan speed until Thermal Management is activated | 1. Turning system on and verifying that the CFE and MPB fans are at maximum speed. 2. Verifying that after the Thermal Management is activated, the CFE, MPB and BE fans are at minimum speed.
42 | 07 - CFE - System shut down due to overheat - Upper Limit(For Trident) | 1. Setting one of the sensors readings, associated with the CFE fan, above "Upper Limit" for more than 6 minutes.
43 | 10 - CFE - System shutdown due to overheat - Critical sensor above Valid Upper limit(For Trident) | 1. Setting a Critical sensor reading, associated with the CFE fan, above "Valid Upper Limit" for more than 6 minutes.
44 | 03 - VENUE_FE Periodic Monitoring | Verifying that the system periodically monitors the following: 1. FE status indications 2. FE voltages 3. FE currents 4. FE temperatures 5. FE fan speed
45 | 04 - C-BEP Periodic Monitoring (Trident) | Verifying that the system periodically monitors the following: 1. C-BEP status indications 2. C-BEP voltages 3. C-BEP temperatures 4. C-BEP fan speed
46 | 06 - MPB Periodic Monitoring | Verifying that the system periodically monitors the following: 1. MPB status indications 2. MPB Power 3. MPB temperatures 4. MPB fan speed
47 | 03 - Fan mulfunction | 1. Disconnecting one of the systems fans and verifying that a failure alert message is displayed to the user. 2. Disconnecting one of the systems fans and verifying that a warning is written to the system log.
48 | 11 - CFE - System shutdown due to overheat - Critical sensor below Valild Lower limit | 1. Setting a Critical sensor reading, associated with the CFE fan, below "Valid Lower Limit" for more than 6 minutes.
49 | 07 - MPB - System shut down due to overheat - Upper Limit | 1. Setting the "MPB.MPB_TEMP_SENSOR" sensor reading above "Upper Limit" for more than 6 minutes. 2. Setting the "BAT1.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 3. Setting the "BAT2.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 4. Setting the "BAT3.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 5. Setting the "BAT4.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 6. Setting the "BAT5.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 7. Setting the "BAT6.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes.
50 | 10 - MPB - System shutdown due to overheat - Critical sensor above Valid Upper limit | 1. Setting the "MPB.MPB_TEMP_SENSOR" sensor reading above "Valid Upper Limit" for more than 6 minutes. 2. Setting the "BAT1.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 3. Setting the "BAT2.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 4. Setting the "BAT3.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 5. Setting the "BAT4.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 6. Setting the "BAT5.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 7. Setting the "BAT6.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes.
51 | 11 - MPB - System shutdown due to overheat - Critical sensor below Valild Lower limit | 1. Setting the "MPB.MPB_TEMP_SENSOR" sensor reading below "Valid Lower Limit" for more than 6 minutes. 2. Setting the "BAT1.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 3. Setting the "BAT2.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 4. Setting the "BAT3.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 5. Setting the "BAT4.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 6. Setting the "BAT5.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 7. Setting the "BAT6.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes.
52 | 04 - BE - System shutdown due to overheat - Upper Limit | 1. Setting the BEP.TEMP_CPU sensor reading above"Upper Limit" for more than 6 minutes. 2. Setting the BEP.TEMP_SYSTEM sensor reading above "Upper Limit" for more than 6 minutes.
53 | 06 - BE - System shutdown due to overheat - Valid Limit | 1. Setting one of the BE temperature sensors readings above "Valid Upper Limit" for more than 6 minutes. 2. Setting one of the BE temperature sensors readings below "Valid Lower Limit" for more than 6 minutes.
54 | 01 - Temperature sensors reading while in Standby | 1. Verifying that while the system was in Standby, temperature readings from the batteries and MPB were written to the SOM log.
55 | 02 - Temperature sensors reading while in Shutdown | 1. Verifying that while the system was in Shutdown, temperature readings from the batteries and MPB were written to the SOM log.

## Battery Set
| Step | Test Description | Expected Results
Test ID: | 5735
Test Name: | 02 - Battery life threshold
Purpose: | Verifying that Battery Health indication appears according to spec.
Tested Subject: | 1. Verifyting that when the capacity of a full-charged battery is below 50% of a new battery capacity, then the Battery Health indication appears. 2. Verifying that when the capacity of a full-charged battery is above 70% of a new battery capacity, then the Battery Health indication does not appear.
Pre-Conditions: | 1. System is ON 2. Debug Dialog (Alt+F3) is enabled. 3. Two batteries packs are connected to the system 4. See comments
Attachments: | NA
Comments: | 1. Simulating BatteryHealth value shall be done as follows: a) While App is on, press Alt+F3 b) Open IM c) Set Param/Event Name = AcqHkMonCfg d) Object ID: Value = BAT1.BatteryHealth¦SimEnable e) Data:Value = true f) Press "SetParam" button g) Object ID: Value = BAT1.BatteryHealth¦Value h) Data: Value = [Desired Value] i) Press "SetParam" button 2. In order to restore an item value to its actual value, apply the following: a) Object ID: Value = BAT1.BatteryHealth¦SimEnable b) Data: Value = false
Requirement Reference: | 821, 8809
Test | Step 1 | Simulate the Battery Health value to 0.49, as written in the "Commnets" sections > Wait 2 minutes. | A message appears permanently on the display, indicating that the battery needs to be replaced.
| Step 2 | Restore the Battery Health value to its actual value, as written in the "Comments" section > Restart the system > Simulate the Battery Health value of battery 1 to 0.71 > Wait 2 minutes. | A message, indicating that the battery needs to be replaced, does not appear.
Test ID: | 6253
Test Name: | 07 - MPB - System shut down due to overheat - Upper Limit(For Trident)
Purpose: | Verifying that if one of the MPB associated sensor readings is above "Upper Limit" for more than 6 minutes, the system will shut down according to spec.
Tested Subject: | Setting one of the sensors associated with the MPB fan, above "Upper Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on. 2. All the MPB fan associated sensors readings are within the allowed limits.
Attachments: | NA
Comments: | The Monitoring Item names of the sensors assiciated with the MPB fan are: 1. MPB.MPB_TEMP_SENSOR 2. BAT1.BatteryTemperature 3. BAT2.BatteryTemperature 4. BAT3.BatteryTemperature 5. BAT4.BatteryTemperature 6. BAT5.BatteryTemperature 7. BAT6.BatteryTemperature The sensors limits can be found in here: %target_root%\resources\acquisition\housekeeping\Mon_BAT_FuncRev1A-POC-Trident.xml %target_root%\resources\acquisition\housekeeping\Mon_MPB_FuncRev1A.xml Simulating a sensor overheating is done like so: 1. Go to %target_root%\resources\idunn\setup\debug.res and uncomment ShowDebugDialog 2. start application and open debug dialog window (Alt+F3) 3. Tap IM... and under Param/Event Name type AcqHkMonCfg {Case Sensitive - Pay attention!} 4. Object ID --> Value= SensorName¦SimEnable (For example: MPB.MPB_TEMP_SENSOR¦SimEnable ) 5. Data --> Value= true 6. Tap SetParam button 7. Object ID --> Value= SensorName¦Value (For example: MPB.MPB_TEMP_SENSOR¦Value ) 8. Data --> Value= NumericValue (For example: 80 ) 9. Tap SetParam button 10. Wait for 6 minutes
Requirement Reference: | 2254
Test | Step 1 | Set one one of the sensors, associated with the MPB fan, above"Upper Limit" and below "Valid Upper Limit" > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
Test ID: | 6254
Test Name: | 04 - BE - System shutdown due to overheat - Upper Limit(For Trident)
Purpose: | Verifying that if one of the BE temperature sensors readings are above "Upper Limit" for more than 6 minutes, the system will shut down and a message shall appear to the user.
Tested Subject: | 1. Setting one of the sensors associated with the BEP fan, above"Upper Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on. 2. All the BE fan associated sensors readings are within the allowed limits.
Attachments: | NA
Comments: | The Monitoring Item names of the sensors assiciated with the BEP fan are: 1. BEP.TEMP_CPU 2. BEP.TEMP_SYSTEM
Requirement Reference: | 2254, 9958
Test | Step 1 | Set one one of the sensors, associated with the BEP fan, above"Upper Limit" and below "Valid Upper Limit" > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
Test ID: | 6283
Test Name: | 01 - Battery capacity display
Purpose: | Verifying that the scanner display remaining battery capacity
Tested Subject: | 1.Verifying that the system has a battery indicator 2.Verifying that the user is notified when the remaining system battery capacity is less than 10 min.
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in live B mode 3. The battery capacity is a little above 10 min 4. AC cable is disconnected
Attachments: | NA
Comments: | N\A
Requirement Reference: | 2357, 2359
Test | Step 1 | Verify that the scanner has a battery indicator on the display showing remaining battery capacity. | The scanner has a battery indicator on the display showing remaining battery capacity.
| Step 2 | Scan until the battery capacity descends below 10 min | A user notification appears when the remaining system battery capacity is below 10 min
Test ID: | 6311
Test Name: | 05 - Error indicator of disabled battery charge
Purpose: | Verifying that a special error indication appears when AC plug is connected to power but charging of battery is disabled.
Tested Subject: | Verifying that a special error indication appears when AC plug is connected to power but charging of battery is disabled
Pre-Conditions: | 1. AC plug is connected to power 2. The system is ON 3. Debug Dialog (Alt+F3) is enabled.
Attachments: | NA
Comments: | 1. Simulating Charging error indication shall be done as follows: a) While App is on, press Alt+F3 b) Open IM c) Set Param/Event Name = AcqHkMonCfg d) Object ID: Value = BAT1.ChargeError¦SimEnable e) Data:Value = true f) Press "SetParam" button g) Object ID: Value = BAT1.ChargeError¦Value h) Data: Value = 1 i) Press "SetParam" button 2. In order to restore an item value to its actual value, apply the following: a) Object ID: Value = BAT1.ChargeError¦SimEnable b) Data: Value = false
Requirement Reference: | 2683
Test | Step 1 | Simulate the ChargeError value of battery to "1", as written in the "Commnets" sections > Wait 2 minutes. | Error indication appears.
Test ID: | 6343
Test Name: | 02 - Low battery pop up message
Purpose: | Verifying that a pop up messsage about low battery capacity apperas according to spec.
Tested Subject: | 1. Verifying that a pop up messsage about low battery capacity apperas when a remaining time is 10 min. 2. Verifying that a pop up messsage about low battery capacity re-apperas every 2 minutes after tapping ok on the message. 3. Verifying that the system automaticaly trigers "End Exam" and turns Off the system when remaining time is 2 min.
Pre-Conditions: | 1. Battery scanning time is above 10 min. 2. No power cable connected to the system 3. The system is ON 4. A probe is connected o the system 5. Patient's exam is active 6. The system is in B-mode
Attachments: | NA
Comments: | NA
Requirement Reference: | 9636
Test | Step 1 | Wait till the system remaining time is 10 min | A pop up message appears that notifies about low battery capacity
| Step 2 | Tap on "OK" to close the pop up message | 1. Pop up message is closed 2. The system remains in active patient's exam
| Step 3 | Tap on 'Store'. | An image or cineloop appears in the Clipboard.
| Step 4 | Wait for 2 minutes | A pop up message re-appears that notifies about low battery capacity
| Step 5 | Tap on "OK" > wait till pop up messages appear and close them till the remaining time reaches 2 minutes | 1. The system automaticaly triggers "End Exam" 2. System turns Off
| Step 6 | Connect the system to the power > Turn on the system | The last exam, which was active before the system shut down, was saved, including images/cineloops.
Test ID: | 5742
Test Name: | 06 - Special indication when battery scanning time is less than 10 min
Purpose: | Verifying that battery indicator (auxiliary display) has a special indication when remaining scan time is less than 10 min.
Tested Subject: | Verifying that battery indicator (auxiliary display) has a special indication when remaining scan time is less than 10 min.
Pre-Conditions: | 1. The system is ON 2. A probe is connected to the system. 3. The system is in live B- mode. 4. Battery scan time is more than 10 min. 5. AC cable is disconnected
Attachments: | NA
Comments: | NA
Requirement Reference: | 825
Test | Step 1 | Wait untill the remaining scan time, indicated on the battery indicator, is 10 min. | Venue Go - LED bar is blinking RED Venue - scan time screen blinking RED

## Burning Set
| Step | Test Description | Expected Results
Test ID: | 11731
Test Name: | CFE GEN2 FW burning via gficonsole tool
Purpose: | Verifying it is possible to update FE Monitoring FW version via gficonsole tool
Tested Subject: | The ability to update FE Monitoring FW version via gficonsole tool
Pre-Conditions: | 1. Technical Plug is connected 2. System is in maitenance mode (Windows) 3. FE Monitoring FW version is burned on the system (REV 1.6)
Attachments: | NA
Comments: | NA
Requirement Reference: | NA
Test | Step 1 | Start EchoLoader application. Go to 'Settings' > 'Config' > 'About' > 'Firmware Version' | FE Monitoring FW version is displayed as REV 1.6
| Step 2 | Go to %target_root%\resources\acquisition\HAL\TL_FEU_Actions_CFE2.xml. Change expectedFEMonitorDeviceVersion parameter to 1.7 | expectedFEMonitorDeviceVersion parameter is changed to 1.7
| Step 3 | Open ScLogWindow and CMD window.In CMD - type gficonsole -f and press ENTER | FE Monitoring FW version burning started.Progress is displayed in ScLogWindow
| Step 4 | Wait until burning process is completed | A message is displayed indicating the burning process completed successfully
| Step 5 | Shutdown the system and disconnect all power supplies.Reconnect AC cable or batteries and restart the system | Application is loaded.Correct FE Monitor version appears under System > Consfig >About >Firmware Version .REV 1.6 appears, and 'expected REV 1.7' appears under the comment column.
| Step 6 | Go to %target_root%\resources\acquisition\HAL\TL_FEU_Actions_CFE2.xml and restore expectedFEMonitorDeviceVersion to the original value. | expectedFEMonitorDeviceVersion parameter is changed.
Test ID: | 11735
Test Name: | FW reading from EEPROM
Purpose: | Verifying system reads correcty all boards FW from EEPROM
Tested Subject: | Verifying each board details are read and presented correctly under Config
Pre-Conditions: | NA
Attachments: | NA
Comments: | NA
Requirement Reference: | NA
Test | Step 1 | Open CMD and type gfitest -i | VPD tool is opened.
| Step 2 | Select each one of the boards from the list and for each board tap 'Read VPD'. Document boards details as displayed in VPD tool. | Each boards details are read and displayed in tool.
| Step 3 | Exit VPD tool and start application. Go to Settings > Config > About > HW Version Verify for each board that displayed information matches details displayed in VPD tool | All details displayed in Config are identical to details displayed in VPD tool

## Display Set
| Step | Test Description | Expected Results
Test ID: | 5747
Test Name: | 01 - Auto adjustment of display brightness
Purpose: | Verifying that the monitor's Brightness is auto adjusted
Tested Subject: | Auto adjustment of the monitor's brightness .
Pre-Conditions: | 1. System is ON 2. Auto Sense is enabled (in LCD settings) 3. The system is in Scanning screen
Attachments: | NA
Comments: | NA
Requirement Reference: | 892, 891
Test | Step 1 | Cover light sensors of the screen | Brightness of the screen dims
| Step 2 | Uncover light sensors on the screen | Brightness of the screen is brighten
| Step 3 | In the scan controls decrease or increase brighthness manualy | 1. The system brighthness is changed 2. Auto Sense is swithched to disabled

## Labeling Set
| Step | Test Description | Expected Results
Test ID: | 6286
Test Name: | 04 - MI limitation for Opthalmic preset
Purpose: | Verifying that the Mechanical Index (MI) is not more than 0.23 for optalmic presets
Tested Subject: | 1. Verifying that the Mechanical Index (MI) is less than 0.23 for optalmic preset in 12L-RS probe 2. Verifying that the Mechanical Index (MI) is less than 0.23 for optalmic preset in 12Ln-RS probe 3. Verifying that the Mechanical Index (MI) is less than 0.23 for optalmic preset in 3Sc-RS probe 4. Verifying that the Mechanical Index (MI) is less than 0.23 for optalmic preset in 9L-RS probe
Pre-Conditions: | 1. 9L-RS, 12L-RS, 12Ln-RS, 3Sc-RS probes are available 2. The system is in Home screen
Attachments: | NA
Comments: | NA
Requirement Reference: | 2413
Test | Step 1 | Activate "Ophthalmic" preset in "12L" probe | 1. The system is in live B mode 2. 12L shortcut button is highlighted 3. Opthalmic preset is selected 4. Acoustic output (AO) is 100 % 5. Mechanical index (MI) is less than 0.23
| Step 2 | Activate "Ophthalmic" preset in "12Ln" probe | 1. The system is in live B mode 2. 12Ln shortcut button is highlighted 3. Opthalmic preset is selected 4. Acoustic output (AO) is 100 % 5. Mechanical index (MI) is less than 0.23
| Step 3 | Activate "Ophthalmic" preset in "3Sc" Probe | 1. The system is in live B mode 2. 3Sc shortcut button is highlighted 3. Opthalmic preset is selected 4. Acoustic output (AO) is 100 % 5. Mechanical index (MI) is less than 0.23
| Step 4 | Activate "Ophthalmic" preset in "9L" probe | 1. The system is in live B mode 2. 9L shortcut button is highlighted 3. Opthalmic preset is selected 4. Acoustic output (AO) is 100 % 5. Mechanical index (MI) is less than 0.23
Test ID: | 6339
Test Name: | 07 - The factory name of preset Ophthalmic
Purpose: | Verifying that the name of the factory preset for Ophthalmic scanning is Ophthalmic
Tested Subject: | 1. The name of the Ophthalmic factory preset on L12n-RS probe 2. The name of the Ophthalmic factory preset on 12L-RS probe 3. The name of the Ophthalmic factory preset on 3Sc-RS probe
Pre-Conditions: | 1. 12Ln-RS, 12L-RS and 3Sc-RS probes are connected 2. System is in 'Home' screen
Attachments: | NA
Comments: | N/A
Requirement Reference: | 9627
Test | Step 1 | Select L12n-RS probe and Ophthalmic factory preset | 1. Scanning screen is displayed 2. The L12n-RS probe is the current probe 3. Ophthalmic factory preset is selected and displayed under the probe name 4. Ophthalmic factory preset is also displayed in the title bar
| Step 2 | From scanning screen tap on 12L-RS probe | 1. The system remains in scanning screen 2. The probe is switched to 12L-RS 3. Ophthalmic factory preset is selected and displayed under the probe name 4. Ophthalmic factory preset is also displayed in the title bar
| Step 3 | From scanning screen tap on 3Sc-RS probe | 1. The system remains in scanning screen 2. The probe is switched to 3Sc-RS 3. Ophthalmic factory preset is selected and displayed under the probe name 4. Ophthalmic factory preset is also displayed in the title bar
Test ID: | 7402
Test Name: | 09- Labeling - Service Manuals
Purpose: | Verifying that Reference Manual contains the required information for the service personnel
Tested Subject: | The Service Manual provides information for service personnel regarding: 1. FRU replacements 2. Installations 3. Maintenance 4. Quality checks
Pre-Conditions: | NA
Attachments: | NA
Comments: | NA
Requirement Reference: | 2160, 2161, 2162, 2163, 2164, 2165, 2166, 2167, 9808, 2169, 2170, 2391, 2392, 2393, 2394, 2395, 2396, 2397, 2400, 9228, 2095, 2020, 2168
Test | Step 1 | Open the "Cockpit Replacement Procedure" section. | The Service Manual providea a warning and instructions how to replace the cockpit in order to avoid any injury during this process.
| Step 2 | Open the "System Modules- Replacement Procedures" section. | There shall be a Basic Service Manual that includes information for service providers which includes FRU replacements, installations, maintenance and quality checks.(Chapter 10)
| Step 3 | Open the "Average Setup Time " section. | Installation (from unpacking to scanning, exluding networking, connectivity configuration) shall not take more than 30 minutes.
| Step 4 | Verify that: Proprietary Service Manual includes information for service personnel regarding FRU replacements,installations, maintenance and quality checks. | Verify that: Proprietary Service Manual includes information for service personnel regarding FRU replacements ,installations, maintenance and quality checks.(Chapter 8)
| Step 5 | Verify that: System has a proprietary service manual | System has a proprietary service manual
| Step 6 | Verify that: Proprietary service manual includes a chapter for Introduction | Proprietary service manual includes a chapter for Introduction
| Step 7 | Verify that: Proprietary service manual includes a chapter for Site Preparations | Proprietary service manual includes a chapter for Site Preparations
| Step 8 | Verify that: Proprietary service manual includes a chapter for System Setup | Proprietary service manual includes a chapter for System Setup
| Step 9 | Verify that: Proprietary service manual includes a chapter for General Procedures and Functional Checks | Proprietary service manual includes a chapter for General Procedures and Functional Checks
| Step 10 | Verify that: Proprietary service manual includes a chapter for Components and Function (Theory) | Proprietary service manual includes a chapter for Components and Function (Theory)
| Step 11 | Verify that: Proprietary service manual includes a chapter for Service Adjustments | Proprietary service manual includes a chapter for Service Adjustments
| Step 12 | Verify that: Proprietary service manual includes a chapter for Diagnostics/Troubleshooting | Proprietary service manual includes a chapter for Diagnostics/Troubleshooting
| Step 13 | Verify that: Proprietary service manual includes a chapter for Replacement Procedures | Proprietary service manual includes a chapter for Replacement Procedures
| Step 14 | Verify that: Proprietary service manual includes a chapter for Parts Renewal | Proprietary service manual includes a chapter for Parts Renewal
| Step 15 | Verify that: Proprietary service manual includes a chapter for Care and Maintenance | Proprietary service manual includes a chapter for Care and Maintenance
| Step 16 | Verify that: The Proprietary Service Manual includes FRU instructions and information. | Service Manual includes FRU instructions and information.
| Step 17 | Verify that: Service Manual includes System Setup instructions, including Installation Wizard | Service Manual includes installation instructions.
| Step 18 | Verify that Service Manual contains information regrding "Insite" | Servicee Manual contains information regarding "InSite"
| Step 19 | Verify that: Service Manual includes set up instructions. | Service Manual includes set up instructions.
| Step 20 | Verify that: Service Manual includes User and FE maintenance instructions | Service Manual includes User and FE maintenance instructions (Chapter 10)
| Step 21 | Verify that: Service Manual includes troubleshooting aids. | Service Manual includes troubleshooting aids.
| Step 22 | Verify that: Service Manual contains Caution/Warning instruction to take extreme caution when service personnel are handling the equipment | Service Manual contains Caution/Warning instruction to take extreme caution when service personnel are handling the equipment
| Step 23 | Open the "Electrical Safety Tests" section. | Service Manual includes instructions to do System safety test after repair or installation of safety related modules or grounding.

## PreventiveActions Set
| Step | Test Description | Expected Results
Test ID: | 6246
Test Name: | 01 - Preventation of scanning in case of over voltage
Purpose: | Verifying that the system monitors voltage and stops scanning in case of over voltage
Tested Subject: | 1. Verifying that the a probe is scanning with in limit transmit voltages 2. Verifying that the probe stops scanning in case of transmission is over voltage
Pre-Conditions: | 1. The system is ON 2. Service dongle is connected to the system 3. A 3Sc-RS probe is connected to the system 4. The system is in Maintenance
Attachments: | NA
Comments: | After the test undo all file changes [System type] in test can be 'Trident' or 'Hawaii'
Requirement Reference: | 2222
Test | Step 1 | Double tap on "Echoloader" shortcut | 1. Echoloader application is loaded 2. 3Sc-RS probe is in live B mode
| Step 2 | Close Echoloader application and exit to Windows > Navigate to %target_root%\resources\acquisition\Probes\ | "3Sc-RS_[System type].xml" file exists
| Step 3 | Open "3Sc-RS_[System type].xml" file with 'Notepad' application > change "txVoltageAllowedDeviation" to -0.5 and "VoltageerrorLimit2" to 0 > save file | "3Sc-RS_[System type].xml" file was modifyed
| Step 4 | Double tap on "Echoloader" shortcut from desktop | 1. Echoloader application is loaded 2. System is not scanning
Test ID: | 6247
Test Name: | 01 - Scanning relation to FE clock frequency
Purpose: | Verifying that the Image degradates due to HW error
Tested Subject: | Verifying that the Image degradates due to HW error
Pre-Conditions: | 1. A linear probe is connected and selected. 2. The system is in Windows screen 3. ScLogWindow is running (with "Always on top" option)
Attachments: | NA
Comments: | NA
Requirement Reference: | 2244
Test | Step 1 | Run "EchoLoader" application | ScLogWindow contains the following message: "Clock frequency check module, initial check passed"
| Step 2 | In Windows activate command prompt ( CMD ) | Command Prompt window appears
| Step 3 | Type "cd /d %target_root%\bin64" and tap on "Enter Button" | The current working directory has been changed
| Step 4 | Run GFIConsole.exe | "CFI>" command line appears
| Step 5 | Type: dm 0xc0200672 > tap "Enter" | GFIConsole window resault is: - 0xc0200672 -0x007f
| Step 6 | In GFIconsole window tipe: wm 0xc0200670 0x3 > tap "Enter" | 1. Major setup error window pops 2. Scan stopped 3. SclogWindow contains the following: "FE INTF Error" "FE_INTF CLOCK_FREQ_ERR" status register data Clock accurracy info.
| Step 7 | Restart "EchoLoader" application > restart GFIconsole application | 1. EchoLoader" application is running 2. GFIconsole is running
| Step 8 | In GFIconsole window type: wm 0xc0200670 0x4 > Tap "Enter" | 1. ScLogWindow shows: Clock frequency ¦FE INTF Error¦FE_INTF CLOCK_FREQ_ERR ¦status register data¦ Clock accurracy info 2. Major setup error window pops. 3. Scan stopped
| Step 9 | Restart "EchoLoader" application > restart GFIconsole | 1. "EchoLoader" application running 2. GFIconsole is running
| Step 10 | In GFIconsole window type: wm 0xc0200670 0x8 > Tap "Enter" | 1. ScLogWindow shows: "Clock frequency¦FE INTF Error¦FE_INTF CLOCK_FREQ_ERR¦status register data¦Clock accurracy info 2. Major setup error window pops. 3. Scan stopped
Test ID: | 6340
Test Name: | 01 - Periodic system shutdown mechanism
Purpose: | Verifying that the system has Periodic shutdown mechanism
Tested Subject: | 1. Verifying that user is able to postpone periodic shutdown 2. Verifying that periodic shutdown mechanism is done upon user aproval
Pre-Conditions: | 1. A probe is connected to the system 2. The next parameters in %target_root%\resources\ idunn\ setup\ globalconfig.res file - RebootTimeoutDaysValue to 0 - RebootTimeoutWarning to 0:0:1 - RebootTimeoutRepeatedWarning to 0:0:1 3. The system is in B live mode
Attachments: | NA
Comments: | After running the test undo the modifications to %target_root%\resources\ idunn\ setup\ globalconfig.res file
Requirement Reference: | 9629
Test | Step 1 | Wait at least for 1 minute > press on "On/Off" button > tap on "Standby" button | 1. A window with 'Shutdown' and 'Postpone' buttons appear 2. A message in a window recommends to reboot the system to maintain performance
| Step 2 | Tap on "Postpone" button | The system enters into a standby mode
| Step 3 | Press on "On/Off" button, after the system enters into a live scanning mode , press on "On/Off" button again and tap on "Standby" | A message in a window recommends to reboot the system to maintain performance
| Step 4 | Tap on "Postpone" button to Enter Standby mode. Press on "On/Off" button . | The system is in live B mode
| Step 5 | Wait at least 1 minute > press on "On/Off" button > tap on "Standby" button | 1. A window with 'Shutdown' and 'Postpone' buttons appear 2. A message in a window recommends to reboot the system to maintain performance
| Step 6 | Tap on "Shutdown" button | The system shuts down

## ProbesFailureHandling Set
| Step | Test Description | Expected Results
Test ID: | 6243
Test Name: | 01 - Probe ID comparison in the system
Purpose: | Verifying that the software compares the probe ID found in file with the ID obtained from probe
Tested Subject: | Verifying that supported probe is not supported if Probe ID in system file was changed
Pre-Conditions: | 1. The system is ON 2. Technical plug is connected to the system 3. 3Sc-RS probe is connected to the system 4. The system is in Patient screen
Attachments: | NA
Comments: | After the test undo all file changes
Requirement Reference: | 2218
Test | Step 1 | Verify that 3Sc-RS probe appears in the probe selection menu | 3Sc-RS probe appears in the probe selection menu
| Step 2 | Tap on any probe's preset | The system is in B live mode
| Step 3 | Enter to Maintanace mode (windows) > Navigate to %target_root%\resources\acquisition\Probes > open "3Sc-RS.xml " file with a notepad application | Probe ID Exists in the file
| Step 4 | Change probe ID field to 000 > save file > enter to Echoloader application | Echoloader application is loaded
| Step 5 | tap on "Patient" button | 1. 3Sc-RS probe does not appear in the probe selection menu 2. Scan is not possible with this probe
Test ID: | 6244
Test Name: | 02 - Probe not supported if probe ID is duplicated
Purpose: | Verifying that supported probe is not supported ifprobe ID is duplicated in the system
Tested Subject: | Verifying that supported probe is not supported ifprobe ID is duplicated in the system
Pre-Conditions: | 1. The system is ON 2. The system is in Maintenance 3. 9L-RS and 3Sc-RS probes are available for test
Attachments: | NA
Comments: | After the test undo all file changes
Requirement Reference: | 2219
Test | Step 1 | Navigate to %target_root%\resources\acquisition\Probes | "3Sc-RS.xml" and "9L-RS.xml" files exist
| Step 2 | Open "3Sc-RS.xml" and "9L-RS.xml" files with 'Notepad' application | "3Sc-RS.xml" and "9L-RS.xml" files are opened
| Step 3 | In "9L-RS.xml" file change probe ID field to the same number as in "3Sc-RS.xml" file | Probe ID in "9L-RS.xml" same as in "3Sc-RS.xml" file
| Step 4 | Save "9L-RS.xml" file > enter to Echoloader application | Echoloader application is loaded
| Step 5 | Connect 9L-RS and 3Sc-RS probes to the system > tap on "Home" button | 1. 3Sc-RS probe does not appear in the probe selection menu 2. 9L-RS probe does not appear in the probe selection menu 3. Scan is not possible 4. Probe slots are empty on the screen

## SwWarningsNotifications Set
| Step | Test Description | Expected Results
Test ID: | 6256
Test Name: | 01 - Presence of Thermal and Mechanical indexes
Purpose: | Verifying that Thermal and Mechanical indexes are presented
Tested Subject: | Verifying that Thermal and Mechanical indexes are presented
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in 2D scanning mode 3. Additonal equipment: Tissue phantom, Gammex 403 GS LE or similar
Attachments: | NA
Comments: | N\A
Requirement Reference: | 2261
Test | Step 1 | Verify that: 1. Thermal index appears on title bar 2. Mechanical index appears on the title bar | 1. Thermal index appears on title bar 2. Mechanical index appears on the title bar
Test ID: | 6257
Test Name: | 02 - Correct presentation of geometry data
Purpose: | Verifying that geometry data represented correctly
Tested Subject: | Verifying that geometry data represented correctly by using tissue phantom
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in 2D scanning mode 3. Additonal equipment: Tissue phantom, Gammex 403 GS LE or similar
Attachments: | NA
Comments: | N\A
Requirement Reference: | 2262
Test | Step 1 | Place a probe on tissue phantom | 1. Clear image appears 2. Geometry of tissue phantom is correct 3. Grey scale targets are not deformed
Test ID: | 6258
Test Name: | 03 - Correct representation of colour coded data
Purpose: | Verifying that colour coded data represented correctly
Tested Subject: | Verifying that colour coded data represented correctly by using flow phantom
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in 2D scanning mode 3. Flow phantom set to pulse mode with preset 3 4. Additonal equipment: Tissue phantom, Gammex 403 GS LE or similar
Attachments: | NA
Comments: | N\A
Requirement Reference: | 2263
Test | Step 1 | Tap on "CFM" button > place probe on flow fantom > locate the flow in the tube | 1. Coded colour data appears inside CFM ROI 2. Coded colour data appears inside the tube only
Test ID: | 6261
Test Name: | 06 - Confirmation of patient data assignment
Purpose: | Verifying that the user must confirm intended assignment of unassigned patient
Tested Subject: | Verifying that the user must confirm intended assignment of unassigned patient
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in Live B-Mode with unassigned patient 3. Additonal equipment: Tissue phantom, Gammex 403 GS LE or similar
Attachments: | NA
Comments: | N\A
Requirement Reference: | 2267
Test | Step 1 | Tap on "Store" button | A cine appears in clipboard
| Step 2 | Tap on "Patient" button | "Resume exam" and "End Exam" buttons appear
| Step 3 | Tap on "End Exam" | Massage appears "Patient information was not entered"
| Step 4 | Tap on "Enter patient info" button > fill Patient ID field > tap on "End Exam" > Tap on "Yes" | 1. "End exam" button is inactive 2. "Begin Exam" button appears 3. Patient ID field is empty
| Step 5 | Tap on "Local Archive" tab button | A new created Patient ID exists in the list
Test ID: | 6262
Test Name: | 07 - Includement specific risks per tool in user manual
Purpose: | Verifying that the user manual includs specific risks per tool according to spec.
Tested Subject: | Verifying that the user manual includs specific risks for the following automatic M&A tools: 1. Auto-VTI 2. Auto-IVC 3. B-Lines
Pre-Conditions: | 1. The system is ON 2. User manual is opened 3. Additonal equipment: Tissue phantom
Attachments: | NA
Comments: | N\A
Requirement Reference: | 2268
Test | Step 1 | Verify that the user manual includs specific risks for the following automatic M&A tools in the user manual: 1. Auto-VTI 2. Auto-IVC 3. B-Lines | Specific risks for the following automatic M&A tools in the user manual appears : 1. Auto-VTI 2. Auto-IVC 3. B-Lines
Test ID: | 6264
Test Name: | 11 - M&A graphics position
Purpose: | Verifying that the M&A is performed, remove the measurements from the image upon changing the layout of a combined image
Tested Subject: | Verifying that the M&A is performed, remove the measurements from the image upon changing the layout of a combined image
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in PW scaning mode 3. Additional equipment: Flow phantom
Attachments: | NA
Comments: | N\A
Requirement Reference: | 2275
Test | Step 1 | Tap on 'Freeze' > Tap on "Measure" button > perform any measuremet | A measurement appears on the screen
| Step 2 | Review the available scan controls | The button "Layout" is unavailable
| Step 3 | Tap 'Store' and recall stored image from clipboard. | Measure graphics shown on the scanned image
| Step 4 | Review the available scan controls | The button "Layout" is available
| Step 5 | Tap on 'Layout' button | The measurement remove from the image.
Test ID: | 6265
Test Name: | 12 - Discontinuities in time motion display
Purpose: | Verifying that moving TM-cursor results in recordings with discontinuites are indicated on the display
Tested Subject: | Verifying that moving TM-cursor results in recordings with discontinuites are indicated on the display
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in PW scaning mode 3. Additional equipment: Flow phantom, Gammex OPTIMIZER 1425A or similar.
Attachments: | NA
Comments: | N\A
Requirement Reference: | 2276
Test | Step 1 | Pan PW cursor on 2D image > tap on "Freeze" button | Discontinuities indicated by lightning sign
Test ID: | 6266
Test Name: | 13 - Colour flow presentation
Purpose: | Verifying that colour flow processing on flow phanton represented correctly
Tested Subject: | Verifying that colour flow processing on flow phanton represented correctly
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in live B mode
Attachments: | NA
Comments: | Additional equipment : Flow phantom that configured to : a. Pulsed flow b. Preset 3
Requirement Reference: | 2277
Test | Step 1 | Tap on "CFM" button | CFM ROI appears
| Step 2 | Put a probe on flow phantom > find a tube with a flow | 1. Flow is represented by a colour 2. Colour stays inside a tube
Test ID: | 6267
Test Name: | 14 - M&A identifiers
Purpose: | Verifying that M&A identifires appear aside each measurement in resault screen
Tested Subject: | Verifying that M&A identifires appear aside each measurement in resault screen
Pre-Conditions: | 1. A probe is connected to the system 2. The system is in live B-mode 3. Additonal equipment: Tissue phantom, Gammex 403 GS LE or similar
Attachments: | NA
Comments: | NA
Requirement Reference: | 2281
Test | Step 1 | Tap on "Cardiac" category > tap on "Measure" button tap on "Dimension" tool > perform few measurements | 1. Measurements result window appears 2. Near each result there is an identifier belonging to each measurement
Test ID: | 6268
Test Name: | 16 - Confirmation of data deletion
Purpose: | Verifying of confirmation of data deletion
Tested Subject: | 1. Tapping on "No" in deletion confirmation window 2. Tapping on "Yes" in deletion confirmation window
Pre-Conditions: | 1. The system is in 'Archive management' screen 2. A patient exists in 'Archive management'
Attachments: | NA
Comments: | NA
Requirement Reference: | 2286
Test | Step 1 | Tap on a patient > tap on " Delete Selected" button | Message appears "The selected item(s) will be deleted. Please confirm"
| Step 2 | Tap on "Cancel" button | Patient stays in Archive manager
| Step 3 | Tap on a patient > tap on " Delete Selected" button | Message appears "The selected item(s) will be deleted. Please confirm"
| Step 4 | Tap on "Ok" button | A patient disappears from Archive manager
Test ID: | 6269
Test Name: | 28 - Real-time Clock
Purpose: | Verifying that the system provides the real-time clock in the scanning modes
Tested Subject: | 1. The real-time clock indicates the current time in the Live scanning modes 2. The real-time clock indicates the current time in the Freeze state
Pre-Conditions: | 1. A probe is connected to the system 2. The system in Live B mode 3. 'Date and Time' settings adjusted according to the region time zone
Attachments: | NA
Comments: | NA
Requirement Reference: | 2287
Test | Step 1 | Compare the time displayed by the real-time clock on the title bar with the current time | 1. The real-clock displays the current time 2. Time in the real-clock runs
| Step 2 | Tap on 'Freeze' button | 1.'Freeze' button turns blue 2. The time in the clock stopps running and displays the time of the freeze
Test ID: | 6270
Test Name: | 17 - Patient data replacement message
Purpose: | Verifying that a warning massage that a patient data will be replaced with the backup copy appears
Tested Subject: | Verifying that a warning massage that a patient data will be replaced with the backup copy appears
Pre-Conditions: | 1. 'Patient Archive' backup was performed (Config > Admin > Backup) 2. All of the patients were deleted from the Archive Management (After backup) 3. The system is in 'Config'screen
Attachments: | NA
Comments: | N/A
Requirement Reference: | 2289
Test | Step 1 | In 'Admin' > 'Restore' > Mark the 'Patient Archive' checkbox > Tap on 'Restore Now' | A warning message with indication that archive is being replaced with the backup copy appears
| Step 2 | Tap 'Yes' > 'OK' > Go to 'Archive Management' | Patients list is restored
Test ID: | 6271
Test Name: | 27 - Keep or replace message when storing to media
Purpose: | Verifying that the message asking to keep or replace already existing data on media while storing to media appears
Tested Subject: | Verifying that the message asking to keep or replace already existing data on media while storing to media appears
Pre-Conditions: | 1. The system is ON 2. A USB memstick connected to the system 3. The same exam exists in USB memstick as RAW Data 4. The system is in "Archive Management" page 5. The dataflow is configured with DICOM Raw data
Attachments: | NA
Comments: | NA
Requirement Reference: | 2290
Test | Step 1 | Select the same patient in "Archive management" that exists on USB memstick > Tap on "Send selected" > Select USB memstick > tap on "OK" button | Message asking to keep or replace already exiting file on the media appears
| Step 2 | Tap on 'Cancel transfer' > Select the same patient in "Archive management" that exists on USB memstick > Tap on "Send selected" > Select Dicom USB memstick > tap on "OK" button | The file has been sent to destination without warning message
Test ID: | 6272
Test Name: | 18 - No space on media report
Purpose: | Verifying that copying data to a full media not performed and a failure reported
Tested Subject: | Verifying that copying images to a full media not performed and failure reported
Pre-Conditions: | 1. A patient with at least two exams and each exam has cine loops exists in the system 2. The system is in Archive Management screen 3. USB memstick connected to the system (See comments)
Attachments: | NA
Comments: | USB memstick is full (not enough to copy a cine loops) .
Requirement Reference: | 2291
Test | Step 1 | Tap on a patient with exams and cineloops > tap on "Send selected" button > select "DICOM USB memstick" in 'TO' field > tap on "OK" button | 1. Copy operation not performed 2. A message notifing that the disk is full appears
Test ID: | 6273
Test Name: | 19 - Network monitoring and patient demographic backup
Purpose: | Verifying that network is monitored and patient demographic data is backed up on local archive
Tested Subject: | 1. Verifying that network is monitored 2. Verifying that patient demographic data is backed up on the system 3. Verifying that patient data is backed up on the Local archive
Pre-Conditions: | 1. The system is configured to store patient data to DICOM server with "Worklist/Local Archive - DICOM Server/Int. HD" dataflow 2. The system is connected to the network with LAN cable 3. A new patient with demographic data is loaded 4. A probe is connected to the system 5. The system is in live B-mode
Attachments: | NA
Comments: | NA
Requirement Reference: | 2292
Test | Step 1 | Verify that Network monitoring icon shows that network is connected | Network monitoring icon shows that network is connected
| Step 2 | Disconnect a Network Cable from the system | Network monitoring icon shows that network is not available
| Step 3 | Tap on "Store" button > tap on "End Exam" > tap on "Local Archive" tab | A patient which exam was recently closed , appears in Local Archive
| Step 4 | In 'Local Archive' tap on a patient which exam was recently closed and continue the exam | 1. The system is in B live mode 2. Stored cine appears in a clipboard
| Step 5 | Tap on "Home" button | In 'Current Patient' tab, patient's demographics data appears
Test ID: | 6274
Test Name: | 20 - SW version and HW configuration on a configuration info screen
Purpose: | Verifying that SW and HW configuration appears on configuration information screen
Tested Subject: | 1. Verifying that SW version appears on configuration information screen 2. Verifying that HW configuration appears on configuration information screen
Pre-Conditions: | 1. The system is ON 2. The system is in SW version in configuration information screen (under 'About')
Attachments: | NA
Comments: | NA
Requirement Reference: | 2293
Test | Step 1 | Tap on 'System Version' button | SW version appears on configuration information screen
| Step 2 | Tap on "HW version" button | 1. Systems HW configuration list appears 2. Revision of each HW appears
Test ID: | 6275
Test Name: | 21 - Rejection of image in spooler
Purpose: | Verifying that storage job spooler notifies about images that were rejected
Tested Subject: | Verifying that storage job spooler notifies about images that were rejected
Pre-Conditions: | 1. The system is ON 2. The system is connected to the network 3. A probe is connected to thesystem 4. The configured connectivity dataflow is 'Local Archive - Int HD/DICOM' (Marked as 'Default') 5. The configured IP address of the DICOM server is wrong 6. The port number in TCP/IP configuration and Dicom port number are identical 7. A patient is loaded 8. The system is in live B- mode
Attachments: | NA
Comments: | N/A
Requirement Reference: | 2294
Test | Step 1 | Tap on "Store" button 3 times> Tap on 'End Exam' >Tap on Alt + S | 1. Storage job spooler opens 2. The 3 images that have been stored are displayed as rejected in the spooler list (Status = 'Failed')
Test ID: | 6276
Test Name: | 22 - Probe visualisation by the system
Purpose: | Verifying that the system provides probe visualisation during operation and during selection
Tested Subject: | a) Verifying that the system provides probe menu during selection b) Verifying that the system provides probe display during operation
Pre-Conditions: | 1. The system is ON 2. 3Sc-RS, C1-5-RS, 12L-RS probes connected to the system 3. The system is in Home screen
Attachments: | NA
Comments: | NA
Requirement Reference: | 2296
Test | Step 1 | Verify that 3Sc-RS, C1-5-RS, 12L-RS probes and probes menues appear on the screen | 3Sc-RS, C1-5-RS, 12L-RS probes and probes menues appear on the screen
| Step 2 | Tap on any probes application | 1. The system is in B scanning mode 2. Probes name and probes application appears in toolbar area
Test ID: | 6277
Test Name: | 23 - Preventation of scanning if HW not valid
Purpose: | Verifying that scanning is prevented if HW cofiguration validation failed
Tested Subject: | Verifying that scanning is prevented if HW cofiguration validation failed
Pre-Conditions: | 1. The system is in Windows 2. Probe is connected to the system 3. One not valid revision HW board installed in the system (see comments)
Attachments: | NA
Comments: | Write in windows serch GfiTest -i (or in cmd) Run gfitest -i to change HW configuration (Functional Revesion) Select CFE> tap edit Change 'Functional Revision' to 5A (other than exists, write down original value) tap on 'SET CRC' Exit After test is complete: Write in windows serch GfiTest -i Run gfitest -i to change HW configuration (Functional Revsion) Select CFE > tap edit Change 'Functional Revision' back to 1A (back to original value) tap on 'SET CRC' Exit
Requirement Reference: | 2264, 2297
Test | Step 1 | Run EchoLoder application | 1.The system doesn't enter to EchoLoder application. 2. Error message appears
Test ID: | 6278
Test Name: | 24 - Control for Image Quality optimisation
Purpose: | Verifying that the system provides a control to activate test-pattern for Image Quality optimisation
Tested Subject: | Verifying that the system provides a control to activate test-pattern for Image Quality optimisation
Pre-Conditions: | 1. The system is ON 2. A probe is connected to the system 3. The system is in live B-mode 4. Additonal equipment: Tissue phantom, Gammex 403 GS LE or similar
Attachments: | NA
Comments: | NA
Requirement Reference: | 2299
Test | Step 1 | Tap on Settings > LCD > Test image button | Test Pattern' button appears
| Step 2 | Tap on 'Test Pattern' widget > Tap on '+' / '-' | Screen color pattern is changed
Test ID: | 6562
Test Name: | 02 - message about light level
Purpose: | Verifying that the system provides a message to the user thatambient light level exceeds the recommended level and image quality may be less than optimal
Tested Subject: | message appears when light level exceeds the recommended level
Pre-Conditions: | 1. System is ON 2. Auto Sense is enabled (in LCD settings) 3. The system is in Scanning screen 4. Addditional equipment: Flash light
Attachments: | NA
Comments: | NA
Requirement Reference: | 9891
Test | Step 1 | Bring excessive light to the light sensor of the system | Message to the user appears about ambient light level exceeds the recommended level (image quality may be less than optimal)

## ThermalControl Set
| Step | Test Description | Expected Results
Test ID: | 6249
Test Name: | 02 - Temperature and Fan speed Periodic Monitoring
Purpose: | Verifying that the Thermal Management module is monitoring Temperatures and Fans speed according to spec.
Tested Subject: | 1. Verifying that the log contains periodic readings of temperatures sensors from the following components: a) FE b) BEP c) Batteries 2. Verifying that the log contains periodic readings of the systems fans.
Pre-Conditions: | System is on, in Maintenance mode.
Attachments: | NA
Comments: | NA
Requirement Reference: | 2253, 2250
Test | Step 1 | Exit to windows > Go to D:\Log > Open file logfile-TempMon.txt | The file contains readings, every 2 minutes, of the following: 1. FE temperature sensors 2. BEP temperature sensors 3. Batteries temperature sensors 4. FE fan speed 5. BEP fan speed
Test ID: | 6251
Test Name: | 01 - Fan speed until Thermal Management is activated
Purpose: | Verifying that all fans, except the BE fan, are initiated with maximum speed until the Thermal Management is activated. Verifying that in room temperature, all fans are at minumum speed, once the Thermal Management is activated.
Tested Subject: | 1. Turning system on and verifying that the CFE and MPB fans are at maximum speed. 2. Verifying that after the Thermal Management is activated, the CFE, MPB and BE fans are at minimum speed.
Pre-Conditions: | 1. System is off. 2. System is placed in room temperature
Attachments:
Comments: | If a step refers to log files, one should enter Maintenance mode. For instructions see attached file.
Requirement Reference: | 2252
Test | Step 1 | Turn system on > Wait 6 minutes > Open file logfile-TempMon.txt > Observe the 1st log row, related to the current boot up. | 1. The CFE fan is at maximum speed. 2. The MPB fan is at maximum speed.
| Step 2 | Observe the 2nd log row, related to the current boot up. | 1. The CFE fan is at minimum speed. 2. The MPB fan is at minimum speed. 3. The BE fan is at minimum speed.
Test ID: | 6252
Test Name: | 07 - CFE - System shut down due to overheat - Upper Limit(For Trident)
Purpose: | Verifying that if one of the CFE associated sensor reading is above "Upper Limit" for more than 6 minutes, the system will shut down according to spec.
Tested Subject: | 1. Setting one of the sensors readings, associated with the CFE fan, above "Upper Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on. 2. All the CFE fan associated sensors readings are within the allowed limits.
Attachments: | NA
Comments: | The Monitoring Item names of the sensors assiciated with the CFE fan are: 1. MPB.FE_TEMP_SENSOR 2. SSCI.Temperature 3. DACQ.Temperature 4. TRX1.Temperature 5. TRX2.Temperature To simulate temperatures: Enable debug window and tap alt+f3 Tap IM in Param/event name write: AcqHkMonCfg 1.under Object Id: under "Value": enter <sensor name>¦SimEnable (e.g: TRX1.Temperature¦SimEnable) 2.under Data: in Value type: true (to enable or false to disable) and tap SetParam 3. repeat step 1 with <sensor name>¦Value 4. repeat step 2 with the desired value to simulate the files with the limits are at %target_root%\resources\acquisition\Housekeeping \resources\acquisition\Housekeeping\Mon_CFE2_FuncRev1A \resources\acquisition\Housekeeping\Mon_CFEB_FuncRev1B
Requirement Reference: | 2254
Test | Step 1 | Set one of the sensors readings, associated with the CFE fan, above Upper Limit and below Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
Test ID: | 6255
Test Name: | 10 - CFE - System shutdown due to overheat - Critical sensor above Valid Upper limit(For Trident)
Purpose: | Verifying that if a Critical sensor reading, associated with the CFE fan, is outside the Valid upper limit for more than 6 minutes, the system will shut down according to spec.
Tested Subject: | 1. Setting a Critical sensor reading, associated with the CFE fan, above "Valid Upper Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on.
Attachments: | NA
Comments: | To simulate temperatures: Enable debug window and tap alt+f3 Tap IM in Param/event name write: AcqHkMonCfg 1.under Object Id: under "Value": enter <sensor name>¦SimEnable (e.g: TRX1.Temperature¦SimEnable) 2.under Data: in Value type: true (to enable or false to disable) and tap SetParam 3. repeat step 1 with <sensor name>¦Value 4. repeat step 2 with the desired value to simulate the limits are located in: %target_root%\resources\acquisition\Housekeeping \resources\acquisition\Housekeeping\Mon_CFE2_FuncRev1A \resources\acquisition\Housekeeping\Mon_CFEB_FuncRev1B CFE fan sensors MPB.FE_TEMP_SENSOR SSCI.Temperature DACQ.Temperature TRX1.Temperature TRX2.Temperature
Requirement Reference: | 2256
Test | Step 1 | Set one of the Critical sensor readings, associated with the CFE fan, above Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
Test ID: | 6123
Test Name: | 03 - VENUE_FE Periodic Monitoring
Purpose: | Verifying that the system periodically monitors the FE according to spec.
Tested Subject: | Verifying that the system periodically monitors the following: 1. FE status indications 2. FE voltages 3. FE currents 4. FE temperatures 5. FE fan speed
Pre-Conditions: | 1. System is ON, in maintenance mode.
Attachments: | NA
Comments: | NA
Requirement Reference: | 1812, 10237, 10239
Test | Step 1 | Exit to windows > Go to D:\Log > Open files: logfile-DACQStatusMon.txt logfile-SSCIStatusMon.txt | The files contain status indications of the DACQ and SSCI every 2 mintues
| Step 2 | Open file: logfile-DCVoltMon.txt | The file contain FE voltages and currents readings every 2 mintues
| Step 3 | logfile-TempMon.txt | The file contain FE temperatures and FE fan speed readings every 2 mintues (Inlet and outlet fans speepds)
Test ID: | 6124
Test Name: | 04 - C-BEP Periodic Monitoring (Trident)
Purpose: | Verifying that the system periodically monitors the C-BEP according to spec.
Tested Subject: | Verifying that the system periodically monitors the following: 1. C-BEP status indications 2. C-BEP voltages 3. C-BEP temperatures 4. C-BEP fan speed
Pre-Conditions: | 1. System is ON, in maintenance mode.
Attachments: | NA
Comments: | NA
Requirement Reference: | 1813
Test | Step 1 | Exit to windows > Go to D:\Log | 1. There is a file containing entires, every 2 minutes, of C-BEP status indications. 2. There is a file containing entries, every 2 mintues, of C-BEP voltages. 3. There is a file containing entries, every 2 mintues, of C-BEP temperatures. 4. There is a file containing entries, every 2 mintues, of C-BEP fan speed.
Test ID: | 8000
Test Name: | 06 - MPB Periodic Monitoring
Purpose: | Verifying that the system periodically monitors the MPB according to spec.
Tested Subject: | Verifying that the system periodically monitors the following: 1. MPB status indications 2. MPB Power 3. MPB temperatures 4. MPB fan speed
Pre-Conditions: | 1. System is ON, in maintenance mode. 2. The system is in "Home" screen
Attachments: | NA
Comments: | NA
Requirement Reference: | 1814
Test | Step 2 | Open keyboard and tap on Alt+D > tap on "Save and export" > Exit to windows > Go to D:\Log\ Zip log file \Log\Logs | 1. There is a file containing entires, every 2 minutes, of MPB status indications. (pslog-general.log) 2. There is a file containing entries, every 2 mintues, of MPB power. (pslog-MPB_VOLTAGE_MON.log )
| Step 3 | D:/Log | 1. There is a file containing entries, every 2 mintues, of MPB temperatures. (logfile-TempMon.txt ) 2. There is a file containing entries, every 2 mintues, of MPB fan (logfile-TempMon.txt )
Test ID: | 1618
Test Name: | 03 - Fan mulfunction
Purpose: | Verifying that in case of a fan malfunction, the Thermal Management module will alert according to spec.
Tested Subject: | 1. Disconnecting one of the systems fans and verifying that a failure alert message is displayed to the user. 2. Disconnecting one of the systems fans and verifying that a warning is written to the system log.
Pre-Conditions: | 1. System is on. 2. One or more fans are disconnected.
Attachments: | NA
Comments: | If a step refers to log files, one should enter Maintenance mode > Exit to Windows > Go to D:\Log
Requirement Reference: | 7068
Test | Step 1 | Verify that a failure alert message is displayed to the user. | A failure alert message is displayed to the user.
| Step 2 | Open file logfile.txt | A warning, related to fan malfunction, is written in the log.
Test ID: | 1692
Test Name: | 11 - CFE - System shutdown due to overheat - Critical sensor below Valild Lower limit
Purpose: | Verifying that if a Critical sensor reading, associated with the CFE fan, is outside the Valid Lower limit for more than 6 minutes, the system will shut down according to spec.
Tested Subject: | 1. Setting a Critical sensor reading, associated with the CFE fan, below "Valid Lower Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on.
Attachments: | NA
Comments: | CFE fan sensors MPB.FE_TEMP_SENSOR SSCI.Temperature DACQ.Temperature TRX1.Temperature TRX2.Temperature Enable debug window and tap alt+f3 Tap IM in Param/event name write: AcqHkMonCfg 1.under Object Id: under "Value": enter <sensor name>¦SimEnable (e.g: TRX1.Temperature¦SimEnable) 2.under Data: in Value type: true (to enable or false to disable) and tap SetParam 3. repeat step 1 with <sensor name>¦Value 4. repeat step 2 with the desired value to simulate %target_root%\resources\acquisition\Housekeeping the limits are located in one of the files: \resources\acquisition\Housekeeping\Mon_CFE2_FuncRev1A \resources\acquisition\Housekeeping\Mon_CFEB_FuncRev1B
Requirement Reference: | 7079
Test | Step 1 | Set one of the Critical sensor readings, associated with the CFE fan, below Valid Lower Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown 2. System shuts down.
Test ID: | 1739
Test Name: | 07 - MPB - System shut down due to overheat - Upper Limit
Purpose: | Verifying that if one of the MPB associated sensor readings is above "Upper Limit" for more than 6 minutes, the system will shut down according to spec.
Tested Subject: | 1. Setting the "MPB.MPB_TEMP_SENSOR" sensor reading above "Upper Limit" for more than 6 minutes. 2. Setting the "BAT1.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 3. Setting the "BAT2.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 4. Setting the "BAT3.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 5. Setting the "BAT4.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 6. Setting the "BAT5.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes. 7. Setting the "BAT6.BatteryTemperature" sensor reading above "Upper Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on. 2. All the MPB fan associated sensors readings are below "Upper Limit" and above "Lower Limit".
Attachments: | NA
Comments: | To simulate temperatures: Enable debug window and tap alt+f3 Tap IM in Param/event name write: AcqHkMonCfg 1.under Object Id: under "Value": enter <sensor name>¦SimEnable (e.g: TRX1.Temperature¦SimEnable) 2.under Data: in Value type: true (to enable or false to disable) and tap SetParam 3. repeat step 1 with <sensor name>¦Value 4. repeat step 2 with the desired value to simulate The limits are located in: bat-Mon_BAT_FuncRev1A-POC-Trident \resources\acquisition\Housekeeping\Mon_CFE2_FuncRev1A \resources\acquisition\Housekeeping\Mon_MPB_FuncRev1A
Requirement Reference: | 7088, 2258
Test | Step 1 | Set the "MPB.MPB_TEMP_SENSOR" sensor reading above Upper Limit and below Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 2 | Turn the system on > Restore the "MPB.MPB_TEMP_SENSOR" sensor reading below Upper Limit and above Lower limit > Set the "BAT1.BatteryTemperature" above Upper Limit and below Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 3 | Turn the system on > Restore the "BAT1.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT2.BatteryTemperature" above Upper Limit and below Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 4 | Turn the system on > Restore the "BAT2.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT3.BatteryTemperature" above Upper Limit and below Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 5 | Turn the system on > Restore the "BAT3.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT4.BatteryTemperature" above Upper Limit and below Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 6 | Turn the system on > Restore the "BAT4.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT5.BatteryTemperature" above Upper Limit and below Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 7 | Turn the system on > Restore the "BAT5.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT6.BatteryTemperature" above Upper Limit and below Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
Test ID: | 1741
Test Name: | 10 - MPB - System shutdown due to overheat - Critical sensor above Valid Upper limit
Purpose: | Verifying that if a Critical sensor reading, associated with the MPB fan, is outside the Valid upper limit for more than 6 minutes, the system will shut down according to spec.
Tested Subject: | 1. Setting the "MPB.MPB_TEMP_SENSOR" sensor reading above "Valid Upper Limit" for more than 6 minutes. 2. Setting the "BAT1.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 3. Setting the "BAT2.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 4. Setting the "BAT3.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 5. Setting the "BAT4.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 6. Setting the "BAT5.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes. 7. Setting the "BAT6.BatteryTemperature" sensor reading above "Valid Upper Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on. 2. All the MPB fan associated sensors readings are below "Upper Limit" and above "Lower Limit".
Attachments: | NA
Comments: | To simulate temperatures: Enable debug window and tap alt+f3 Tap IM in Param/event name write: AcqHkMonCfg 1.under Object Id: under "Value": enter <sensor name>¦SimEnable (e.g: TRX1.Temperature¦SimEnable) 2.under Data: in Value type: true (to enable or false to disable) and tap SetParam 3. repeat step 1 with <sensor name>¦Value 4. repeat step 2 with the desired value to simulate The limits are located in: bat-Mon_BAT_FuncRev1A-POC-Trident \resources\acquisition\Housekeeping\Mon_CFE2_FuncRev1A \resources\acquisition\Housekeeping\Mon_MPB_FuncRev1A
Requirement Reference: | 7090
Test | Step 1 | Set the "MPB.MPB_TEMP_SENSOR" sensor reading above Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 2 | Turn the system on > Restore the "MPB.MPB_TEMP_SENSOR" sensor reading below Upper Limit and above Lower limit > Set the "BAT1.BatteryTemperature" above Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 3 | Turn the system on > Restore the "BAT1.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT2.BatteryTemperature" above Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 4 | Turn the system on > Restore the "BAT2.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT3.BatteryTemperature" above Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 5 | Turn the system on > Restore the "BAT3.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT4.BatteryTemperature" above Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 6 | Turn the system on > Restore the "BAT4.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT5.BatteryTemperature" above Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 7 | Turn the system on > Restore the "BAT5.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT6.BatteryTemperature" above Valid Upper Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
Test ID: | 1743
Test Name: | 11 - MPB - System shutdown due to overheat - Critical sensor below Valild Lower limit
Purpose: | Verifying that if a Critical sensor reading, associated with the MPB fan, is outside the Valid Lower limit for more than 6 minutes, the system will shut down according to spec.
Tested Subject: | 1. Setting the "MPB.MPB_TEMP_SENSOR" sensor reading below "Valid Lower Limit" for more than 6 minutes. 2. Setting the "BAT1.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 3. Setting the "BAT2.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 4. Setting the "BAT3.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 5. Setting the "BAT4.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 6. Setting the "BAT5.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes. 7. Setting the "BAT6.BatteryTemperature" sensor reading below "Valid Lower Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on. 2. All the MPB fan associated sensors readings are below "Upper Limit" and above "Lower Limit".
Attachments: | NA
Comments: | To simulate temperatures: Enable debug window and tap alt+f3 Tap IM in Param/event name write: AcqHkMonCfg 1.under Object Id: under "Value": enter <sensor name>¦SimEnable (e.g: TRX1.Temperature¦SimEnable) 2.under Data: in Value type: true (to enable or false to disable) and tap SetParam 3. repeat step 1 with <sensor name>¦Value 4. repeat step 2 with the desired value to simulate The limits are located in: bat-Mon_BAT_FuncRev1A-POC-Trident \resources\acquisition\Housekeeping\Mon_CFE2_FuncRev1A \resources\acquisition\Housekeeping\Mon_MPB_FuncRev1A
Requirement Reference: | 7090
Test | Step 1 | Set the "MPB.MPB_TEMP_SENSOR" sensor reading below Valid Lower Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 2 | Turn the system on > Restore the "MPB.MPB_TEMP_SENSOR" sensor reading below Upper Limit and above Lower limit > Set the "BAT1.BatteryTemperature" below Valid Lower Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 3 | Turn the system on > Restore the "BAT1.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT2.BatteryTemperature" below Valid Lower Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 4 | Turn the system on > Restore the "BAT2.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT3.BatteryTemperature" below Valid Lower Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 5 | Turn the system on > Restore the "BAT3.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT4.BatteryTemperature" below Valid Lower Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 6 | Turn the system on > Restore the "BAT4.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT5.BatteryTemperature" below Valid Lower Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 7 | Turn the system on > Restore the "BAT5.BatteryTemperature" sensor reading below Upper Limit and above Lower limit > Set the "BAT6.BatteryTemperature" below Valid Lower Limit > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
Test ID: | 1793
Test Name: | 04 - BE - System shutdown due to overheat - Upper Limit
Purpose: | Verifying that if one of the BE temperature sensors readings are above "Upper Limit" for more than 6 minutes, the system will shut down and a message shall appear to the user.
Tested Subject: | 1. Setting the BEP.TEMP_CPU sensor reading above"Upper Limit" for more than 6 minutes. 2. Setting the BEP.TEMP_SYSTEM sensor reading above "Upper Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on. 2. All the BE fan associated sensors readings are inside the limits.
Attachments: | NA
Comments: | 1. If a step refers to log files, one should enter Maintenance mode > Exit to Windows > Go to D:\Log To simulate temperatures: Enable debug window and tap alt+f3 Tap IM in Param/event name write: AcqHkMonCfg 1.under Object Id: under "Value": enter <sensor name>¦SimEnable (e.g: TRX1.Temperature¦SimEnable) 2.under Data: in Value type: true (to enable or false to disable) and tap SetParam 3. repeat step 1 with <sensor name>¦Value 4. repeat step 2 with the desired value to simulate The limits are located in: \resources\acquisition\Housekeeping\Mon_BEP_FuncRev1A-POC-Trident
Requirement Reference: | 7097
Test | Step 1 | Set the BEP.TEMP_CPU sensor reading above"Upper Limit" and below "Valid Upper Limit" > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 2 | Turn the system on > Restore the BEP.TEMP_CPU sensor readings inside the limits > Set the BEP.TEMP_SYSTEM sensor reading above"Upper Limit" and below "Valid Upper Limit" > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
Test ID: | 1798
Test Name: | 06 - BE - System shutdown due to overheat - Valid Limit
Purpose: | Verifying that if one of the BE temperature sensors readings is outside "Valid Limits" for more than 6 minutes, the system will shut down and a message shall appear to the user.
Tested Subject: | 1. Setting one of the BE temperature sensors readings above "Valid Upper Limit" for more than 6 minutes. 2. Setting one of the BE temperature sensors readings below "Valid Lower Limit" for more than 6 minutes.
Pre-Conditions: | 1. System is on.
Attachments: | NA
Comments: | 1. If a step refers to log files, one should enter Maintenance mode > Exit to Windows > Go to D:\Log To simulate temperatures: Enable debug window and tap alt+f3 Tap IM in Param/event name write: AcqHkMonCfg 1.under Object Id: under "Value": enter <sensor name>¦SimEnable (e.g: TRX1.Temperature¦SimEnable) 2.under Data: in Value type: true (to enable or false to disable) and tap SetParam 3. repeat step 1 with <sensor name>¦Value 4. repeat step 2 with the desired value to simulate The limits are located in: \resources\acquisition\Housekeeping\Mon_BEP_FuncRev1A-POC-Trident
Requirement Reference: | 7099
Test | Step 1 | Set one of the BE temperature sensors readings above "Valid Upper Limit" > Wait for 6 minutes. | 1. A message appears notifying on system shutdown due to overheat. 2. System shuts down.
| Step 2 | Set one of the BE temperature sensors readings below "Valid Lower Limit" > Wait for 6 minutes. | 1. A message appears notifying on system shutdown 2. System shuts down.
Test ID: | 1761
Test Name: | 01 - Temperature sensors reading while in Standby
Purpose: | Verifying that while the system is in Standby, the Thermal Management is monitoring the temperatures of all connected batteries and MPB.
Tested Subject: | 1. Verifying that while the system was in Standby, temperature readings from the batteries and MPB were written to the SOM log.
Pre-Conditions: | 1. System was in Standby for ~4 minutes.
Attachments: | NA
Comments: | NA
Requirement Reference: | 7092
Test | Step 1 | Turn on the system > Open the SOM log using Alt+D > Export Log > Open folder PSlog. | The log row contains readings, from the time the system was in Standby, of the following temperature sensors:a. All connected batteries.b. MPB
Test ID: | 2432
Test Name: | 02 - Temperature sensors reading while in Shutdown
Purpose: | Verifying that while the system is in Shutdown, the Thermal Management is monitoring the temperatures of all connected batteries and MPB.
Tested Subject: | 1. Verifying that while the system was in Shutdown, temperature readings from the batteries and MPB were written to the SOM log.
Pre-Conditions: | 1. System was in Shutdown for ~4 minutes.
Attachments: | NA
Comments: | NA
Requirement Reference: | 7092
Test | Step 1 | Turn on the system > Open the SOM log using Alt+D > Export Log > Open folder PSlog. | The log row contains readings, from the time the system was in Shutdown, of the following temperature sensors:a. All connected batteries.b. MPB